# Navigating the FDA’s De Novo Program: Incorporating Best Practices from the Beginning

Topic details and speakers background link: https://themedtechconference.com/2025-schedule/?schedule_preset=session%3A3105493%3Adates%3A2025-10-06%3Atags%3Atype-panel-session#session:3105489:dates:2025-10-07:tags:type-panel-session

### Summary

Topic: FDA De Novo interactive review process — lessons, challenges, and strategies for success.

Participants: Josh (moderator), Cody (industry representative), Joshua (FDA representative), Allison, and Laura.

### Key Discussion Points

#### 1. Cody’s Experience with Interactive Review

- Cody described challenges during his company’s De Novo submission to the FDA for a diagnostic device.

- The biggest surprise was a disconnect between the company’s statisticians and the FDA’s biostatistics team regarding study design — particularly around the number of lesion and cancer cases in a reader study.

- Although FDA had reviewed and approved the protocol twice, it later raised major deficiencies.

- To resolve this, the company repeated and expanded studies (eventually completing six clinical and two non-clinical studies).

- They conducted:

  - Multiple reader studies (focusing on melanoma performance).

  - Validation studies in both primary care and dermatology settings.

  - Additional human factors and repeatability studies.

- Cody emphasized flexibility — quickly addressing FDA concerns, running extra studies in parallel, and proactively exceeding requirements.

- He also noted the importance of post-market surveillance for ongoing validation and algorithm improvement.

#### 2. FDA Perspective – Joshua (FDA)

- Explained how post-market surveillance is increasingly used to manage uncertainty in De Novo decisions.

- FDA uses tools like real-world evidence and registries to supplement limited premarket data, particularly for innovative or small-company products.

- Post-market requirements can allow FDA to approve a device with remaining uncertainties while still ensuring safety and effectiveness.

- Communication and collaboration are essential — flexibility and mutual understanding can help both FDA and sponsors reach approval more efficiently.

#### 3. Allison’s Insights

- Supported the value of post-market controls as a way to balance risk and benefit when data are incomplete.

- Emphasized that FDA’s role is to be skeptical — to consider what could go wrong — while industry’s job is to clearly communicate benefits and mitigation strategies.

- Advised companies to “tell their story” and negotiate constructively, using tools like special controls, labeling, and post-market data collection.

#### 4. Diversity, Equity, and Algorithmic Validation

- Cody discussed ensuring equitable performance across skin tones and cancer types, noting FDA’s sensitivity to fairness concerns (referencing issues like pulse oximeters).

- His team aligned with FDA by including diverse cases and integrating post-market monitoring to validate next-generation algorithms.

#### 5. Escalation and Dispute Resolution

- Joshua (FDA) explained the stepwise escalation process for disagreements:

  - Direct discussion with the lead reviewer.

  - Formal justification in the deficiency response.

  - Using the “least burdensome” flag to escalate to an office director.

  - Appeals process if a De Novo is denied.

- Most issues are resolved through early communication rather than escalation.

- Allison added that understanding how FDA writes deficiencies and responding precisely to each “ask” is key.

- Encouraged using submission issue requests and maintaining professionalism — “FDA are people too.”

#### 6. Laura’s Note on the Network of Experts

- Learned about FDA’s Network of Experts program (external advisors) through the FDA ombudsman.

- Although the process was somewhat opaque, she highlighted it as a useful mechanism for resolving technical impasses on effectiveness questions.

#### 7. Closing Takeaways (Josh)

- FDA engagement must be strategic, proactive, and respectful.

- Success depends on thoughtful communication, responsiveness to feedback, and balancing innovation with regulatory rigor.

Overall Theme:
 The De Novo process is collaborative but demanding — success requires flexibility, transparency, and strategic engagement with the FDA throughout premarket, interactive review, and post-market phases.

## Summary: FDA Perspectives on De Novo Classification, Special Controls, and Validation for SaMD

### 1. Classification and Special Controls

- Most devices classified by the FDA are Class II, which rely on special controls to ensure safety and effectiveness (per 21 CFR 860.7 and De Novo Classification Process (Evaluation of Automatic Class III Designation) Guidance, 2021).

- Special controls may include:

  - Biocompatibility and sterility testing for patient-contacting devices.

  - Clinical performance data for novel technologies, including software-driven devices.

  - Post-market data collection requirements that mirror elements of a PMA post-approval study.

- FDA increasingly uses post-market special controls for novel devices—balancing premarket and postmarket data collection to streamline market entry while ensuring long-term safety.

- For Software as a Medical Device (SaMD) and AI/ML-based systems, FDA emphasizes early planning for predetermined change control plans (PCCPs), consistent with FDA’s “Marketing Submission Recommendations for a Predetermined Change Control Plan for AI/ML-Enabled Device Software Functions” (April 2023).

### 2. Importance of Early and Strategic FDA Engagement

- Early interaction with FDA—through Q-submissions (Pre-Subs) or informational meetings—is critical for shaping De Novo and 510(k) strategies.

  - Informational meetings are particularly effective when sponsors don’t yet have specific questions but want FDA reviewers to understand their technology and future plans.

  - Sponsors can present anticipated special controls, draft regulatory language, and future 510(k) pathways during these meetings.

- Establishing FDA reviewers as stakeholders early improves alignment and increases the likelihood of successful review outcomes.

  - Reference: “Requests for Feedback and Meetings for Medical Device Submissions: The Q-Submission Program” (FDA, 2023 update).

### 3. Submission and Review Best Practices

- Day 1–15 post-submission:

  - Ensure all materials are complete for eSTAR submissions, which facilitate organization and transparency.

  - Maintain contact with the lead reviewer to prevent early rejection or technical screening failures.

- Communication practices:

  - Reach out as soon as the reviewer assignment is confirmed.

  - Offer to walk through the file organization and highlight critical data sets.

  - If questions or deficiencies arise, clarify them promptly—unclear responses can cause major review delays.

  - Use rational scientific justification when disagreeing with FDA requests, escalating only when necessary.

### 4. Managing Reviewer Expertise and Clinical Context

- Reviewers may not always be experts in the specific therapeutic area—especially in cross-disciplinary fields like metabolic bone diseases or digital therapeutics.

- Sponsors are encouraged to:

  - Provide subject matter experts (SMEs) during meetings.

  - Encourage reviewers to consult FDA’s Network of Experts Program for additional insight.

  - Reference appropriate SaMD Clinical Evaluation Guidance (IMDRF/SaMD WG/N41:2017) to justify chosen endpoints and effectiveness metrics.

### 5. Communication and Relationship Building

- FDA reviewers emphasize continuous, transparent communication.

  - Developers should treat FDA as a collaborative partner, not merely a regulator.

  - “Tell your story”: When data or testing is questioned, justify why additional information isn’t needed using science-based reasoning and risk management principles (per ISO 14971 and SaMD Risk Categorization (IMDRF/SaMD WG/N12:2014)).

- Effective communication prevents misunderstandings that could lead to deficiency letters or review delays.

### Key FDA Guidance Documents Referenced

- De Novo Classification Process (Evaluation of Automatic Class III Designation) – FDA, 2021

- Q-Submission Program: Requests for Feedback and Meetings for Medical Device Submissions – FDA, 2023

- Marketing Submission Recommendations for a Predetermined Change Control Plan for AI/ML-Enabled Device Software Functions – FDA, 2023

- Software as a Medical Device (SaMD): Clinical Evaluation – IMDRF/SaMD WG/N41:2017

- Software as a Medical Device (SaMD): Risk Categorization and Corresponding Controls – IMDRF/SaMD WG/N12:2014

- ISO 14971:2019 – Application of Risk Management to Medical Devices

### Overall Takeaway

Successful SaMD classification and validation hinge on:

- Proactive, early FDA engagement to define appropriate special controls.

- Clear, justified validation strategies that balance pre- and post-market data.

- Strong communication and collaboration with reviewers throughout the De Novo process.

- Strategic use of FDA guidance frameworks to align SaMD safety, performance, and adaptability requirements with regulatory expectations.