## Global Insights into Implementing Predetermined Change Control Plans

Topic details and speakers background link

https://themedtechconference.com/2025-schedule/?schedule_preset=session%3A3105489%3Adates%3A2025-10-07%3Atags%3Atype-panel-session#session:3105489:dates:2025-10-07:tags:type-panel-session

# Summary: PCCP Panel Session – De Novo AI/ML Medical Devices

### 1. Overview of PCCPs

- Definition: PCCPs are pre-approved plans that describe how a device can evolve over time while maintaining safety and effectiveness.

- Purpose: They allow manufacturers to implement certain modifications post-market without submitting a full new regulatory application.

- Key Benefit: PCCPs encourage innovation, streamline regulatory review, and reduce uncertainty for both manufacturers and regulators.

### 2. FDA Perspective

- FDA has actively supported PCCPs through draft guidance and a centralized cross-office working group since 2024.

- PCCPs apply to all device types, with particular uptake in AI/ML devices and IVDs, due to the need for frequent updates (e.g., new biomarkers, patient populations, or algorithms).

- FDA emphasizes:

  - Early engagement through pre-submissions to discuss proposed changes.

  - Clear, detailed description of intended modifications.

  - PCCPs do not eliminate validation or documentation but shift when testing occurs in the device lifecycle.

- Examples of FDA-approved PCCP changes:

  - Expansion of patient populations for diagnostics.

  - Addition of compatible devices or genetic variants.

  - Updates to antimicrobial susceptibility breakpoints (IVDs).

- Hot tips from FDA:

  - Engage early and deliberately.

  - Provide detailed, high-quality submissions.

  - Be specific about proposed modifications to allow reviewers to assess safety and effectiveness.

  - Recognize that innovative PCCPs may require multiple pre-sub cycles.

### 3. EU & UK Perspectives

- Europe (EU MDR/IVDR):

  - PCCPs are not formally established yet, but discussions are ongoing.

  - There is strong regulatory control due to historical issues (e.g., PIP scandal).

  - IMDRF discussions are informing harmonization across member states.

  - Notified bodies can support structured dialogue but do not provide formal advice like FDA.

- United Kingdom (MHRA):

  - Implementation is slower due to limited resources and evolving policies.

  - Regulatory alignment with FDA and Switzerland is being considered.

  - Predictability is currently low for companies seeking PCCP pathways in the UK.

### 4. Application of PCCPs to De Novo AI/ML Devices

- Why AI/ML fits well:
 AI/ML devices often require iterative updates based on new data, patient populations, or algorithms—making them ideal candidates for PCCPs.

- Regulatory benefits for De Novo submissions:

  - Supports proactive planning for algorithmic or performance updates.

  - Reduces delays in post-market modifications.

  - Ensures consistent safety and effectiveness across updates.

- Best practices:

  - Define specific, high-level modifications upfront.

  - Consider pre-submissions with FDA to validate the PCCP approach.

  - Plan for time investment and discussion cycles; the more innovative the changes, the more iterative engagement is needed.

  - Ensure documentation and validation plans are embedded in the PCCP.

### 5. Tools & Resources

- FDA draft guidance for PCCPs (2024) for all devices.

- IMDRF draft: “Essential Principles and Content of PCCP” – guidance for international harmonization, especially for software updates in medical devices.

- Structured dialogue with notified bodies (EU) for early pathway clarification.

### 6. Key Takeaways

- PCCPs provide a regulatory mechanism for controlled post-market changes, which is crucial for AI/ML devices.

- Early, detailed engagement with FDA or notified bodies is critical for success.

- PCCPs do not reduce validation work; they shift when it occurs, integrating change control into the lifecycle.

- Both EU and UK are exploring PCCPs, but FDA currently offers the most structured and predictable path.

- Companies should strategically plan modifications, pre-sub cycles, and anticipate iterative discussions to maximize PCCP benefits.