PURPOSE

This procedure establishes a standardized, risk‑based process for planning, classifying, approving, and releasing Software as a Medical Device (SaMD) versions as Major, Minor, or Patch releases. It ensures that software changes are adequately assessed, documented, verified, and released in compliance with applicable regulatory and quality requirements.

SCOPE

This procedure applies to all released versions of the SaMD product, including updates to core software, configuration, and integrated off‑the‑shelf (OTS) software components. It covers post‑market updates following the initial regulatory clearance/approval and excludes non‑released internal development builds.

RESPONSIBILITY

| Position |  |
| --- | --- |
| Software Developer | Chief Technology Officer (CTO) |
| Quality Assurance | Regulatory Compliance |
| Machine Learning Developer |  |

REFERENCES

| Document # | Document Title |
| --- | --- |
| IEC 62304:2015 | Medical device software – Software life cycle processes |
| N/A | FDA Guidance Document - Deciding When to Submit a 510(k) for a Software Change to an Existing Device |
| N/A | FDA Webpage - Predetermined Change Control Plans for Machine Learning-Enabled Medical Devices: Guiding Principles |
| N/A | FDA Webpage - Software as a Medical Device (SaMD) |

ASSOCIATED DOCUMENTS

| Document # | Document Title |
| --- | --- |
|  | Change Request Form |
|  | Post-Market Surveillance and Vigilance |
|  | Design History File (DHF) |

DEFINITIONS

| Term | Meaning |
| --- | --- |
| Major Release | A release that introduces significant changes that are likely to require a new or revised regulatory submission (e.g., new 510(k), PMA supplement, EU technical documentation update). Examples include changes to intended use, core algorithms, supported indications, or risk profile. |
| Minor Release | A release that introduces non‑significant changes or updates that are not expected to require a new regulatory submission. Examples include performance optimizations, limited scope feature enhancements, and user interface changes that do not alter intended use or the overall safety and effectiveness profile. |
| Patch Release | A small release primarily intended to correct defects or address urgent cybersecurity vulnerabilities, without introducing new features or functions and without changing the intended use or clinical performance. |
| Change Control Board (CCB) | Cross‑functional team responsible for classification, impact assessment, and approval of software changes (e.g., representatives from Development, Quality, Regulatory Affairs, and Risk Management). |
| Note to File / Regulatory Assessment | Internal documentation capturing the rationale for whether a regulatory submission is or is not required for a given change, aligned with current regulatory guidance on software changes. |

Release Classification

General Principles

Every proposed change to a released SaMD version shall be documented in a Change Request (CR) and submitted to the Change Control Board (CCB) for evaluation.

The CCB shall classify the change as Major, Minor, or Patch according to the criteria below and document the rationale in the CR record.

Major Release

A change shall normally be classified as a Major Release when one or more of the following applies:

- Change to intended use, indications for use, or clinical claims.

- Significant modification of core algorithms or decision logic influencing clinical performance.

- Introduction of new clinical features/functions that may impact diagnosis, treatment, or risk profile.

- Changes that introduce new risks or significantly modify existing risks in a way that could affect safety or effectiveness.

For a Major Release:

- A formal regulatory assessment shall be performed and documented (e.g., Note to File using applicable decision flowcharts).

- Where required, a new or modified regulatory submission shall be prepared and submitted prior to market release.

Minor Release

A change shall normally be classified as a Minor Release when:

- Intended use and clinical claims remain unchanged.

- Changes are limited to performance improvements, usability enhancements, minor feature additions, or non‑significant alterations to existing functionality.

- Risk analysis confirms no increase in overall risk and no new significant hazards.

For a Minor Release:

- A regulatory assessment shall document that a new submission is not required, with justification (e.g., Note to File).

- Design and risk documentation shall be updated as needed to reflect the change.

Patch Release

A change shall normally be classified as a Patch Release when:

- The primary purpose is to correct defects or anomalies, or to address cybersecurity vulnerabilities.

- No new features or functions are introduced.

- No change occurs to intended use, clinical performance, or risk profile beyond reduction of existing risk.

For a Patch Release:

- A concise regulatory assessment shall confirm that the change is limited to bug fixes and/or cybersecurity enhancements and does not require a new submission.

- The Patch Release shall still follow controlled release and testing steps, with scope commensurate to risk.

PROCEDURE

Change Initiation

The originator completes a Change Request, describing:

- Proposed change and rationale,

- Affected components, documents, and configurations, and

- Preliminary assessment of potential impact on safety, performance, cybersecurity, and intended use.

The Change Request is submitted to QA for registration and tracking.

Impact Assessment and Classification

The CCB reviews the Change Request and conducts an impact assessment, including:

- Risk impact analysis (patient safety and cybersecurity),

- Assessment of potential impact on intended use and clinical performance, and

- Assessment of impact on existing regulatory submissions.

The CCB classifies the change as Major, Minor, or Patch and documents the classification rationale and any conditions in the Change Request record.

RA prepares or reviews a regulatory assessment (Note to File or submission plan) and links it to the Change Request.

Planning and Documentation Updates

For all classifications, the Development Lead, with QA and Risk Management, shall:

- Identify all affected documents (e.g., Software Requirements, Design Description, Risk Management File, Verification Plans, SBOM, user documentation), and

- Define required verification and validation activities based on the classification and risk.

For Major Releases, this plan shall align with the company’s product development and design control procedures used for initial release.

Implementation and Testing

Development implements the approved changes in accordance with the Software Development Plan and coding standards.

Testing shall be performed as follows:

- Major Release: Regression testing of relevant previous functionality plus verification and, where required, validation of new or modified functions.

- Minor Release: Risk‑based regression testing plus verification of updated or added functions.

- Patch Release: Focused verification of corrected defects and targeted regression testing of impacted areas, especially security‑related aspects.

All test results are documented and reviewed by QA.

Final Review and Approval

The CCB reviews updated documentation, risk analysis, test evidence, and regulatory assessment.

The CCB confirms the release classification and authorizes the release, if all acceptance criteria are met.

QA ensures that all relevant documents are updated, approved, and stored in the QMS with correct versions and revisions.

Release and Distribution

The released software build is assigned a unique version number according to the defined versioning scheme (e.g., MAJOR.MINOR.PATCH) and linked to the corresponding release classification.

Release notes are prepared, describing:

- Release type (Major, Minor, Patch),

- Summary of changes, and

- Known issues and limitations, if any.

Labeling and user‑facing version information are updated as required.

The release package and associated records (Change Request, regulatory assessment, test reports, risk updates) are archived in the Design History File.

Post-Release Monitoring

Post‑market data (complaints, incidents, cybersecurity alerts, performance metrics) are monitored according to the Post‑Market Surveillance and Vigilance procedures.

Significant issues identified post‑release are evaluated through change control to determine whether a follow‑up Patch, Minor, or Major Release is required.

Versioning Rules

Major Releases increment the major component of the version number (e.g., 1.2.3 → 2.0.0).

Minor Releases increment the minor component (e.g., 1.2.3 → 1.3.0), resetting the patch to zero as needed.

Patch Releases increment the patch component (e.g., 1.2.3 → 1.2.4).

- Note: Document identifiers remain stable; revisions are incremented in accordance with document control procedures for each release.

Records

The following records shall be retained in accordance with record retention procedures:

- Change Requests and CCB minutes.

- Regulatory assessments (Notes to File and, if applicable, submission correspondence).

- Updated design and risk documents.

- Verification and validation reports.

- Release notes and configuration records (including SBOM).

APPROVALS

| Responsible Person | Name/Title |
| --- | --- |
| Document Owner/SME |  |
| Process Owner |  |
| Quality Assurance |  |

REVISION HISTORY

| Rev | Effective Date | Responsible Person | Training Required (Y/N) | Description of Change |
| --- | --- | --- | --- | --- |
| 1.0 | 27 FEB 2026 | Author | Y | New Document |